DC POWER SUPPLY

Ningbo Kaijia Electronic Commerce Co., Ltd.

| Address:Room 609, No. 18 Rikui Road, Jishigang, Haishu District, Ningbo, China | （300W） |
| --- | --- |

Telephone:+86 18757492227

vincent@jesverty.com

USER MANUAL

CONTENTS

SAFETY BRIEF                                    1

SAFETY SYMBOL                                               1

PRODUCT BRIEF                                                2

SPECIFICATION                                                 3

PANEL INSTRUCTION                                       4

WORK REQUIEMENT                                        5

OPERATION INSTRUCTION                             5

CONNECT THE LOAD                                         6

CONSTANT VOLTAGE/CONSTANT                  6

CURRENT CHARACTERISTICS

FUSE REPLACEMENT                                        7

PRODUCT MAINTENANCE                                8

PRODUCT WARRANTY                                      8

PACKING LIST                                                    8

SAFETY BRIEF 	PRODUCT BRIEF

This manual contains important safety instructions that must be followed in the operation and storage environment of theSPS-3005 /3010/6005/12003 series. To ensure your personal safety, and ensures that this product works in the best environment,please read this manual carefully before using.

When you get a brand-new power supply, you need to do the necessary checks to make sure the instrument is working properly.

1. To check whether there are damages caused during	 transportation.

2. To check whether all the accessories are complete.

3.To check whether the output voltage and output current are 	 normal after turning on the device.

If finding out any problems, please contact the merchant immediately.

SAFETY SYMBOL

The safety symbols below will appear in this manual or on the DC power supply.

| ! | Attention High Voltage Grounding |
| --- | --- |

1

The SPS-3005/3010/6005/12003 series of adjustable 
regulated DC power supply designed for use in laboratories, schools and production lines. Both output voltage and 
output current arecontinuously adjustable between 0 and nominal.

The stability and ripple factor of the power supply are very good and have a perfect protection circuit. Can work at full load for a long time. This power supply can be used as both a regulated power supply and a regulated current supply.

2

| SPECIFICATION |  |  |  |  | PANEL INSTRUCTION： |  |  |  |  |  |  | 1 | Output Voltage Display |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1.Switchable DC regulated power supply |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  | 2 | Output Current Display |
| Model Number | SPS-3005 | SPS-3010 | SPS-6005 | SPS-12003 | 4 |  | C.V | V |  |  | 1 | 3 | Output Power Display |
|  |  |  |  |  | A |  |  |  |  |  | 2 | 4 | Voltage Regulator |
| Output Voltage | 0-30V | 0-30V | 0-60V | 0-120V | 5 |  | C.C | W |  |  | 3 | 5 |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  | Current Regulator |
| Output Current | 0-5A | 0-10A | 0-5A | 0-3A |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  | 6 | Voltage Coarse Tuner |
| Input Voltage: AC110V 60Hz |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Working Temperature: 0℃～40℃；Relative Humidity:＜80%RH |  |  |  |  | 8 |  |  | 6 |  |  |  | 7 | Current Coarse Tuner |
| Storage Temperature: -10℃～70℃;Relative Humidity:＜70%RH |  |  |  |  | 10 | V-FINE |  | 5V/2A |  | V-COARSE |  | 8 | Voltage Fine Tuner |
|  |  |  |  |  |  | A-FINE |  |  |  | A-COARSE |  |  |  |
| Constant Voltage State:Voltage stability≤0.1%±3mV |  |  | Voltage stability≤0.2%±3mV |  | 9 |  |  | 7 |  |  |  | 9 | Current Fine Tuner |
| Low voltage:0.2～0.3%±3mV |  |  | Low voltage:0.3～0.4%±3mV |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  | 10 | USB Charging Socket(2A) |
| Load stability≤0.5%±3mV |  |  |  |  | 13 |  |  | 11 |  |  |  | 11 | Positive Polarity(red) |
| Ripple noise≤10mVrms ≤150mVp-p |  |  | Ripple noise≤150mVp-p |  | - |  |  | GND |  | + |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  | 12 | Grounding(green) |
|  |  |  |  |  | 14 |  |  |  | 12 |  |  | 13 | Negative Polarity (black) |
| Constant Current State:Current stability≤0.2%±3mA |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  | 14 | Power Switch |
| Load stability≤0.2%±3mA |  |  |  |  |  |  |  |  |  |  |  |  |  |

Ripple noise≤20mArms

Display Accuracy:0.5%±2digits

Display Resolution:Voltage:00.01V/0.01V  Current:0.001A/0.01A

Product Dimension:220×85×150mm(L×W×H）

| Product Weight:1.4Kg |  |  |  |  | USE ONLY WITH 250V FUSE | 15 | 4 | 15 | Cooling Fan |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Fuse Standard | T6A | T8A | T8A | T8A |  |  |  |  |  |
| The above parameters are measured at an ambient |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | 16 | Power Socket |
| temperature of 25±5°C,relative humidity: <80%RH,and |  |  |  |  |  |  |  |  |  |
| preheated for 30 minutes. The actual parameters will vary |  |  |  |  |  | 16 |  | 17 | Fuse Box |
| slightly. |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  | 17 |  |  |  |

| WORK REQUIEMENT | ! | In actual CV operation, if the load resistance decreases and the |
| --- | --- | --- |
| 1、AC input: Please make sure the input voltage of this product |  | output current increases to the set current value, the power supply |
|  |  | will automatically switch to CC mode.When the load resistance |
|  |  | value continues to decrease, the current will remain at the current |
| AC110V 60Hz |  |  |
|  |  | set value.The voltage is proportionally reduced. At this time, |
| 2、Do not use in an environment where the ambient temperature | Attention | increase the load resistance or increase the current set value to |
|  |  | restore the CV output state. |
| exceeds 40 degrees Celsius. The cooling fan is located at the |  |  |

rear of the device and should have enough space for cooling.

| ! | Incorrect AC voltage input will cause serious damage | CONNECT THE LOAD |
| --- | --- | --- |
|  |  | 1. Rotate the terminal knob by turning it counterclockwise |
|  | to the device. Please make sure the required input |  |
| Warring | voltage value. |  |
|  |  | 2. Insert the load terminal |

3.Turn the terminal knob clockwise

Warring：Incorrect AC voltage input will cause serious damage

to the device. Please make sure the required input voltage value.

OPERATION INSTRUCTION

There are two types of power output modes: constant voltage output (CV) and constant current output (CC). The output mode is determined by the voltage and current values set by the user and the load connected by the user. The output voltage or 
current value of the power supply won't exceed the voltage and current values set by the user. In constant voltage mode, the output voltage value is equal to the user-set voltage value. In constant current mode, the output current value is equal to the user-set current value.

For example: the voltage value is set to 5V and the current value is set to 5A.

Steps

1. Turn on the power switch

2. Adjust the voltage adjustment knob to 5V 
3.Connect the "positive" and "negative"with wires and adjust the current knob to 5A.

4. Disconnect the wire and connect the load to use.

4. Banana plug can be directly inserted into the terminal hole

| ! Attention | Improper connection may result in damage to the power supply and the load connected to the power supply.When connecting the battery load, do not reverse the polarity of the “+” and “-" as this may damage the power supply. |  |
| --- | --- | --- |

CONSTANT VOLTAGE/CONSTANT CURRENT CHARACTERISTICS

The working characteristics of this series of power supplies are constant voltage/constant current automatic conversion type, which can automatically change between constant voltage and constant current states with load changes. The intersection between constant voltage and constant current mode is called conversion point.For example, if the load causes the power supply to operate in a constant voltage mode, a constant voltage is output. As the load increases, the output voltage will remain constant and the output current will increase.When the current value reaches the set current limit value, the power supply will

5 	6

automatically switch to constant current mode. The output

current remains stable and the output voltage decreases 
proportionally as the load increases further.The conversion of constant voltage and constant current is indicated by the LED on the front panel.

CV indicator light is on during constant voltage, CC indicator is on when constant current.

FUSE REPLACEMENT

If the fuse blows, the power supply will stop working.To find and correct the cause of the blown fuse, then replace it with a fuse of the same specification.

Replace the fuse of same

specification,then put the fuse

cabin back.

Remove the power plug first,then open

the fuse cabin according to the

illustration.

Fuse cabin is in the power socket

For effective safety protection, it is only necessary to

4replace the fuse of a specific specification. Before replacing the fuse,the power must be turned off and the

| HIGH VOLTAGE! DANGER! | power cord must be unplugged from the power outlet. |
| --- | --- |

7

PRODUCT MAINTANCE

1.Disconnect the power when the product is not in use.

2.Unplug the power supply before cleaning.

3. Do not use hydrocarbons, chlorides or similar solvents,or use 	 abrasive cleaners.

PRODUCT WARRANTY

1. This product is offered free maintenance service within one year 	 from the date of purchase. Except in the following cases: 
A:Failures caused by improper use, such as improper handling and 	 improper repair, modification or adjustment of the device.

B:Consumable materials are not covered by the warranty.

C:Naturally irresistible disasters such as floods, fires,earthquakes,etc.

2. Maintenance costs are charged for repairs that exceed the warranty  period,and the costs incurred for maintenance are the responsibility  of the user.

PACKING LIST

1. 1x Power Supply

2.1x Power Cord

3. 1x Output Load Cord

4.1x User's Manual

8